## Slide 1
Measurements of neutron response functions for a BC-501A liquid organic scintillator between 5 – 66 MeV at the iThemba LABS high energy neutron beam facility
67th Annual Conference of the South African Institute of Physics “SAIP” 03 – 07 July 2023, University of Zululand, Richards Bay Campus
Kutullo Maibanea, Andy Bufflera, Tanya Huttona, Peane Malekab, Zina Ndabenia,b
a Metrological and Applied Sciences University Research Unit (MeASURe), Department of Physics, University of Cape Town
b iThemba LABS, Faure, Cape Town

## Slide 2
iThemba LABS: k = 200 Separated Sector Cyclotron
National Laboratory
Neutron Vault

## Slide 3
High Energy Neutron Facility at iThemba LABS
Towards an ISO-accredited
high energy neutron metrology facility
at iThemba LABS
continuum well supressed

## Slide 4
  Fluence measurements with a 238U Fission Chamber
  6 tantalum electrodes “90 mm diameter” with fissile materials on both sides - 76 mm sensitive diameter - 99.9823  0.0001 % relative amount of 238U - 443 g/cm2 coating thickness of 238U - 3 kBq activity - 240.931 mg total mass of 238U
Gas flow tubes
Tantalum electrodes
Tantalum window
Incident Neutrons

## Slide 5
The fission reference cross-section (barn) for 238U(n,f)
[IAEA-NDS]
63.5 MeV

## Slide 6
Fluence measurements using the 238U fission chamber at 66 MeV, 8.0 mm natLi, @ 0o, d = 8.00 m
alpha events
fission events
alpha events
Pulse Height spectrum of the 238U FC with the data cut applied to veto counts due to alpha events in the chamber
Corresponding Time of Flight spectra showing the effect of separating the time independent alpha scattering events from the fission events “blue spectrum”
fission events only

## Slide 7
Peak Neutron Fluence
peak neutron fluence [unknown]
number of peak fission events (measured)
fission cross section (known quite well)
number of 238U atoms in FC (known very well)
correction factors for efficiency, neutron absorption and multiplication in the FC, loss of events below threshold, dead time, neutron attenuation in air, beam profile and time resolution (together add about 5% to total uncertainty in fluence)

## Slide 8
Data are recorded in coincidence in list mode using the standard NIM-based acquisition system:
T: time-of-flight S: pulse shape L: pulse height
2” diameter x 4” length, gain-stabilised photomultiplier, neutron-gamma discrimination, time resolution 2 ns “FWHM”
Neutron Spectroscopy: Time of Flight (ToF)
Liquid Scintillator: NE-213 / BC-501A / EJ-301

## Slide 9
Neutrons and gamma-rays produced by a 66 MeV proton beam irradiating a 8.0 mm Li target, measured by a 2” x 4” BC-501A detector at 0o
Pulse height L (MeVee)
Pulse shape S (arbitrary)
neutron-induced events
e
p
d
t
α
gamma ray-induced events
Ralf Nolte
Efficiency calculated with MC codes SCINFUL and MCNPX

## Slide 10
Time-of Flight measurement of neutrons produced by a 66 MeV proton beam irradiating a 8.0 mm natLi target. (BC-501A at 8.00 m from the target at 0o)

## Slide 11
Neutron Spectroscopy: Unfolding
measured neutron pulse height spectrum …
… is unfolded using a
known detector
response matrix …
… to produce the neutron energy spectrum
The full L-T matrix provides the data for selecting nearly monoenergetic pulse height (L) spectra as a function of neutron energy.
UMG-3.3 unfolding package (GRAVEL & MAXED)
Measured neutron pulse height spectrum (channel i)
Detector response matrix (pulse height spectra for monoenergetic neutrons)
Neutron energy spectrum (bin j)

## Slide 12
Measured (tof and unfolded) neutron energy spectra

## Slide 13
Conclusion
Detector response functions were measured at the iThemba LABS fast neutron facility using time-of-flight between 5 - 66 MeV for the reference BC-501A.
There is a good agreement between the measured (tof and unfolded) neutron energy spectra. This provides a validation of the experimentally determined response matrix and of the method used for the unfolding.
The use of unfolding methods and experimentally determined response functions allows neutron spectrometry with BC-501A detector in unknown fields, such as in aircraft and in space.

## Slide 14
…next steps
Further beam characterisations are planned for 100 MeV and 200 MeV, with improved statistics to extend range of applications.
Measurements at 4K binning are planned, which will allow response functions to be determined over the full energy range [5 MeV to 66 MeV and/or above] for every 0.5 MeV.
New or upgraded 238U Fission Chamber detector.

## Slide 15
Thank You…